Micro-ohmmeter

Our company needs Micro-ohmmeter for testing following devices:

- Vacuum circuit breaker

- Power transformer up to 2500 kVA power, 10 kV voltage

- Squirrel cage motor up to 800 kW power, 6 kV voltage

- Current and voltage transformers up to 10 kV voltage

Technical characteristics:

- Measurement range - 1µΩ up to 2500 Ω

- Measurement current – 1mA up to 10 A

- Accuracy class – 0.05%

- Communication output – RS232

- Measurement type – 4 wire kelvin test

- Power supply – rechargeable Battery

Accessories:

- Test leads - 10A kelvin clamps, length ≥3 m – 1 set.

- Caring bag – 1 pc.

- User’s manual on English/Russian language – 1 set

- Data transfer software for PC – 1 pc.

- Communication cable RS232 – 1pc.

- Source cable – 1 pc.