## Sheet1
Angiographic system | Angiographic system
General Information
Quality certificate. CE European, FDA Usa
Release year, no less than 2024
Status new
Technical data
High-tech angiographic system, on which diagnostic and therapeutic angiographic procedures can be performed. With the support of cardio, neurovascular and angioperipheral programs. The system should be able to perform angiography of brain and spinal vessels, endovascular manipulations and operations (aneurysm coiling, balloon-assist, stenting, endovascular thrombectomy, aspiration)
Procedure monitor (ceiling-mounted; movable): Not less than 40 inches, with split screen and control function. Visibility angle not less than 170 degrees
possibility to connect to hospital network and PACS system
Software requirements
QVA- vascular analysis
LVA - analysis of the left ventricle
RVA - right ventricle analysis
QCA-coronary analysis
3D QCA - Analysis of 3D coronaries
3D CT angiography
Roadmap
DSA
Simultaneous imaging of subscripted and non-subscripted images
Determination of the level of stenosis, determination of blood flow
C-arm axis
with ceiling or floor fixing
C-arc depth not less than 92 cm
C-arm axis movement range: (RAO / LAO) not less than + /-105 degrees (CRA / CAU) not less than +45 / -45 degrees The range of SID movement is not less than 90 - 115 cm
The maximum rotation speed is not less than 25 degrees/s
System of automatic protection of patient and personnel from collision of system nodes during movement
Patient table
The possibility of free movement in four directions in the horizontal plane of the table
The range of longitudinal movement of the table is not less than 120 cm
The range of lateral movement of the table is not less than ± 14 cm
Table height movement range of at least 85 - 105 cm from the floor or better
The length of the table is not less than 280 cm
The width of the table is not less than 45 cm
The maximum load of the table more then 190 kg
foot control pedal
Control panel mounted on the table
Upper limbs fixator for radial catheterization
Head fixator for neuro procedures
Doctor's lower body x-ray shield fixed on the table
Function of memorizing positions and activating with one touch
X-ray tube:
Anode heat storage capacity not less than 5.0 MHU
Maximum speed of anodic cooling not less than 7500 HU/S (5.0 kW)
X-ray tube voltage not less than 125 kV
Maximum power not less than 100 kW
The maximum current of the tube is not less than 1000 mA
Maximum uninterrupted input power not less than 3000 W
X-ray beam adjusting collimator with filters
Detector:
The size of the FPD flat detector is not less than 30x30 cm
Detector matrix not less than 1500 × 1500, 14 bit
Active image area not less than 294 mm × 294 mm
The frame rate range of the detector is at least 12-30 FPS
Generator
Power of X-ray high voltage generator not less than 100 kW
Tube current (continuous fluoroscopy) not less than 125 kV, 4 mA
Tube current (pulse fluoroscopy) not less than 120 kV, 200 mA
Additional workstation
Computer with patient registration, data management and archiving forwarding software. Ability to measure, view, edit studies
DICOM 3.0 connectivity, with full import and export functionality.
At least two 19-inch color medical monitors
Ability to record on CD/DVD, USB digital drives
RAM not less than 16 Gb
Memory not less than 1 Tb
Processor frequency not less than 2.8 Ghz
Suitable online UPS for additional workstation
Accessories
Calibration kit
Suspension stand for infusion solutions
Ceiling Mounted Telescopic X-ray Shielding Glass 2.0mm Lead Equivalent
Ceiling-mounted operating lighting LED not less than 60 kLux
Fully equipped safety electrical shield with emergency button
Two-way intercom connection between the procedure and the operator
Online UPS relevant to the power of the angiograph.
Additional requirements:
The supplier must install the system.
With DAP (Dose Area Product)
Remote control and service capability software
User manual of the device
In case of additional request, post-warranty, paid, service possibility for at least 5 years
Staff training
Warranty
Not less than 2 years