## Sheet1
მოთხოვნილი სპეციფიკაცია |  |  | მწარმოებელი ქვეყანა | მწარმოებელი კომპანია | მოდელი | შემოთავაზებული სპეციფიკაცია | ფასი
Holmium Surgical Laser | General | -Holmium laser; -with at least 3 safety glasses; -with fiber service kit; -with foot pedal; -with the time limit removed on fiber; -3 pieces with 800 micron fiber; 3 pieces with 365 micron and 2 pieces with 272 micron fiber (Reusable);
| Power of laser | not less than 60 W
| Frequency | up to 100 Hz or better
| pulse energy range | 0.1 J - 5 J, or better
| pulse duration range | 100-1000 microsecond or better
| wavelenght | not less than 2000 nm
| laser source and class | HO:YAG Class 4
| Control panel | not less than 12 inc; touchscreen
with a guiding beam
| cooling down | Integrated chiller
| Power supply | 220V/50Hz
| Additional requirements | -The seller must provide personal training -Must include user manuals in english
| Certificates | CE or FDA
| Warranty | At least 24 month The warranty must include all necessary services (including preventive maintenance as recommended by the manufacturer)