## Sheet1
Item | Specification | quantiy
ექოსკოპი | Ultrasound System for - cardiovascular diagnostics. System should be of Expert Class, with min. 17' flat panel monitor and workstation, allowing for images archiving on hard disk, CD / DVD; automatic image optimization technique in 2D mode and spectral Doppler, color doppler and continues way doppler should be provided; sectoral probe 2-3,5 MHz, linear probe 5-7,5MHz; microconvex probe 4-7 MHz; TEE probe; Thermal printer with 5rolles of paper should be provided; | 1
ექოსკოპი | field of application: general examinations (abdominal, inc. colon, brest, thiroids); obstetrics and gynecology, System should be of Expert Class, with min. 17' flat panel monitor and workstation, allowing for images archiving on hard disk, CD / DVD; package options to improve the quality of imaging should be provided; software for calculationsshould be provided; Appropriate convex, linear and sectoral probes should be provided; Thermal printer with 5rolles of paper should be provided; | 2
ტრავმატოლოგიური ელექტრო ხერხი |  | 1
კარდიოექოსკოპიური დაჩიკი | თავსებადი LOGIQ-P5-სათვის | 1
კარდიოექოსკოპიური დაჩიკი | თავსებადი medison sonace X6-სთვის | 2
ელექტროკარდიოგრაფი | 12 არხიანი | 3
ხელოვნური სუნთქვის აპარატი | გამოსადეგი ახალშობილების, ბავშვებისა და მოზრდილებისთვის. უნდა ჰქონდეს სუნთქვის შემდეგი რეჟიმები:VCV, PCV, PSV, SIMV(VC/PC) - PS, CPAP, BiPAP, PEEP; პარამეტრების კონტროლი: TV, RR, I:E, PEEP 0-50 cmH20, , FiO2 21-100%, Pressure limit, Triger flow, Triger pressure. აქსესუარები: ნებულაიზერი, დამატენიანებელი. უნდა იყოს მობილური შიდაჰოსპიტალური გადაადგიების საშუალებით. | 5
სანარკოზე აპარატი | გამოსადეგი უნდა იყოს ახალშობილების, ბავშვებისა და მოზრდილებისთვის. უნდა ჰქონდეს სუნთქვის შემდეგი რეჟიმები:VCV, PCV, PSV, SIMV(VC/PC) - PS, CPAP, BiPAP, PEEP; პარამეტრების კონტროლი: TV, RR, I:E, PEEP 0-50 cmH20, , FiO2 21-100%, Pressure limit, Triger flow, Triger pressure. აქსესუარები: ნებულაიზერი, დამატენიანებელი. უნდა იყოს მობილური შიდაჰოსპიტალური გადაადგიების საშუალებით. | 3
კომპიუტერული ტომოგრაფი | 6 სლაისიანი | 1