## Sheet1
Position | specification | Specification Offered | Link | Quantity Required | Quantity Offered | Manufacturing country | Manufacturer | Condition | Year of Manufacturing | Model | Unit Price | Total Price | Currency | unit price in gel | total price in gel | Delivery time | Warranty | Comment | EUR | usd
CT Scanner 128 slice | Advanced CT studies for General Medicine, General Surgery, Vascular surgery, Neurosurgery, Orthopedics and Traumatology, Hepatology-Liver transplantation, Pediatrics. Bone removal; Subtractions for assessment of the blood supply to the brain - CT Angiography, including Liver segmentation, CT Colonoscopy; CT Fluoroscopy – for Lung examinations Power/Generator generator: 100kw Tube helical Unit (Minimum): 7.5 MHU Axial Scan Parameters Scan Speeds: - 0.4, 0.5, 0.6, 0.7, 0.8, 0.9, 1.0 and 2.0 second full scans (360° acquisition); 0.4* is for Pediatric only, Cardiac application: 0.35* - mA capacity: minimum - 10 mA maximum - 850 mA - kVp: 80, 100, 120, 140 Table 70 cm aperture - Vari Speed 0.35 second to 1 second - Remote Tilt +/-30 degrees Image Chain - ASiR radiation dose reduction technology - 64 channel, 128 slice per rotation - 0.35 second rotation - 0.625 mm thin slice images - Two 24” Medical Grade Monitor; Or Two >24” Monitor With QUHD (3800X2600P - (4k)). - 128i overlap reconstruction - English Keyboard, audio contact with main room Image Analysis - Direct MPR – 3 axis recon - Exam Split - Data export and interchange - DICOM worklist - SmartPrep™ contrast monitoring - 3D neuro filters - OptiDose™ – 3D dose modulation - AutomA - Large image series Auto filter & transfer - Dynamic Z-axis tracking - VCT Hi-Power option - High resolution option - Volume Viewer On Console - Low Dose CT lung screening Accessories - Operator chair - Table tray & IV pole - Long cable set - Supports, straps, head holders, pads, wedges, phantoms - Slicker for table and foot pedal For storing and reconstruction of images in Dicom format and 2D and 3D Reconstructions; with possibility to transfer images in DICOM format, to print on Laser or Thermal Printer and to store data on CDR-DVDR; (with appropriate technical characteristics of computer; provider should specify in details) - Injector control - low-dose scanning; - Application Training for medical personal minimum 3 full working days should be provided by supplier; Warranty (Minimum): - full 1 year warranty (from the time when installation protocol will be signed should be provided by Supplier); -5 year - Service warranty. |  |  | 2
CT Scanner 64 slice | Advanced CT studies for General Medicine, General Surgery, Vascular surgery, Neurosurgery, Orthopedics and Traumatology, Hepatology-Liver transplantation, Pediatrics. Bone removal; Subtractions for assessment of the blood supply to the brain - CT Angiography, including Liver segmentation, CT Colonoscopy; CT Fluoroscopy – for Lung examinations Power/Generator generator: 100kw Tube helical Unit (Minimum): 7.5 MHU Axial Scan Parameters Scan Speeds: - 0.4, 0.5, 0.6, 0.7, 0.8, 0.9, 1.0 and 2.0 second full scans (360° acquisition); 0.4* is for Pediatric only, Cardiac application: 0.35* - mA capacity: minimum - 10 mA maximum - 850 mA - kVp: 80, 100, 120, 140 Table 70 cm aperture - Vari Speed 0.35 second to 1 second - Remote Tilt +/-30 degrees Image Chain - ASiR radiation dose reduction technology - 64 channel, 128 slice per rotation - 0.35 second rotation - 0.625 mm thin slice images - Two 24” Medical Grade Monitor; Or Two >24” Monitor With QUHD (3800X2600P - (4k)). - 128i overlap reconstruction - English Keyboard, audio contact with main room Image Analysis - Direct MPR – 3 axis recon - Exam Split - Data export and interchange - DICOM worklist - SmartPrep™ contrast monitoring - 3D neuro filters - OptiDose™ – 3D dose modulation - AutomA - Large image series Auto filter & transfer - Dynamic Z-axis tracking - VCT Hi-Power option - High resolution option - Volume Viewer On Console - Low Dose CT lung screening Accessories - Operator chair - Table tray & IV pole - Long cable set - Supports, straps, head holders, pads, wedges, phantoms - Slicker for table and foot pedal For storing and reconstruction of images in Dicom format and 2D and 3D Reconstructions; with possibility to transfer images in DICOM format, to print on Laser or Thermal Printer and to store data on CDR-DVDR; (with appropriate technical characteristics of computer; provider should specify in details) - Injector control - low-dose scanning; - Application Training for medical personal minimum 3 full working days should be provided by supplier; Warranty (Minimum): - full 1 year warranty (from the time when installation protocol will be signed should be provided by Supplier); -5 year - Service warranty. |  |  | 1

## Sheet2 (hidden)
Condition'
New
Used
Demo
Refurbished