## ფასთა ცხრილი
Item | Specification | Qty | Item Price | Currency | Manufacturer Country | Manufacturer | Model | Delivery Time | Condition | Warranty | Comment | Link/Brochure
სისხლის ხელოვნური მიმოქცევის აპარატი (აეკი)+გათბობა-გაციების სისტემა | აპარატი გვინდა მეორადი ან ქარხნულად აღდგენილი , Stockert S5 + Stockert 3T (მრავალჯერადი ლეიბით და დამაკავშირებელი მილით), Features : 1. Console (movable) on wheels for 5 pumps - four stands + one had pump mounted on a mast. 2. System control panel - integrated control for the entire systems. 3. 4 Peristaltic pumps (roller pumps) - single-head, one (had pump) mounted on a mast. 4. Double-head peristaltic pump – for cardioplegia. 5. Hemodynamic control system - monitoring modules. 5.1. Micro air embolism and liquid level detectors. 5.2. temperature measurement module. 5.3. Timing module - triple stopwatch. 5.3. Timing module - triple stopwatch. 5.4. pressure measurement module - minimum two-channels. 5.5. Cardioplegia module. 5.6. Heat exchanger circuit controlling module. 6. Air-Oxygen gas blender – for example Sechrist or electronic (more expensive). 7. Vacuum assisted venous drainage controller – like this(https://www.perfusion.com/product/digital-vavd-vacuum-regulator-red-0-100mmhg/) or minimum manual venous line occluder . 8. Electrical Venous Occluder (EVO) (https://www.livanova.com/en-GB/Home/Cardiopulmonary/Heart-Lung-Equipment/EVO.aspx). 9. Heater-cooler unit - at least two independent circulation circuits one for oxygenator one for cardioplegia. 10.Optional accessories. 10.1.perfusionist chair. 10.2. LED lamp. 11.Integrated centrifugal pump CP 5 - a drive unit with a control panel, flow sensor, line clamp and manual emergency drive. | 1